Asheville, North Carolina Quality Manual

AxisClinical LLC locations:

- 275 Aiken Road
Asheville, NC  28804  – Manufacturing Plant

- Biltmore Park 
Building 2B  28 Schenck Parkway, Suite 400
Asheville, NC 28803 – Headquarter Offices/Customer Support

- 220 Merrimon Avenue 
Weaverville, NC  28787 – Annex Warehouse and Shipping

- 501 Elk Park Drive 
Woodfin, NC  28804 – Training Center

- 549 Elk Park Drive, Suite 703
Woodfin, NC  28804 – Training Center

- 108 Monticello Road

Weaverville, NC  28787 – Warehouse and Receiving

Website- www.AxisClinical.com

This manual is available via the AxisClinical – ASHEVILLE network.  This document is designed for on-line viewing.  Printed copies, although permitted, are deemed UNCONTROLLED.

QUALITY MANUAL REVISION HISTORY

| Revision | Effective Date | Description Of Change |
| --- | --- | --- |
| AL | 05/06/20 | Changed Management Rep approval from Eddie Reavis to P. John Wojcicki; 1.4 updated Quality Policy; 4.1, page 9 – updated regulator requirements to include “U.S. NRTL and Canadian certification of our products (as applicable) to: UL 471 – Commercial Refrigerators and Freezers UL 499 – Electric Heating Appliances UL/CSA 61010-1 - Electrical Equipment for Measurement, Control and Laboratory Use, General Requirements UL/CSA 61010-2-010 - Particular Requirements for Laboratory Equipment for the Heating of Materials UL/CSA 61010-2-011 - Particular requirements for refrigerating equipment CSA C22.2 No. 120 – Refrigeration equipment CSA C22.2 No. 88 - Construction and Test of Industrial Heating Equipment (EU) 2015/863 – 2015/863 (RoHs 2 amendment) F-gas regulations – (EU) No. 517/2014 Appendix Documentation References: addition of QAP706, Software Development and Maintenance Org Charts – Updated all |
| AM | 1/05/21 | Sections 1.1, 2.0, 4.1: Added references to 93/42/EEC; added missing SOPs to indicate compliance to the Council Directive 93/42/EEC in the Appendix - Documentation References ; Removed reference to ISO9001:2015 |
| AN | 08/06/21 | Page 1 – changed address for 501 Elk Park from “Asheville” to “Woodfin¦ Page 1 – added new Training Center address 549 Elk Park Drive 1.1 changed verbiage to include appropriate number of associates 1.1 revised scope of registration (removing Installation and Service and Technical Support) 2. changed EN ISO14971:2012 to EN ISO 14971:2019 4.1 Added Service and Installation, including Technical Support as an outsourced process 7.5.3 and 7.5.4 Reworded entire section to explain the outsourced process of Service and Installation. 7.5.4 Added paragraph to explain what the Service Repair Depot’s responsibilities are Pages 29 – 33 updated org charts Appendix – documentation references – added QAP405, QAP1410; removed QAP1902 Appendix – responsibilities – removed reference to Technical Support, Service & Installation; changed Process Engineering to Manufacture Engineering |
| AO | 1/26/22 | Section 1.1 revised ISO Scope of Registration to current; Section 4.1 added Korea to “and other applicable regulatory bodies”; Section 4.1 removed reference to Corporate Auditors / added control of outsourced processes and control of AxisClinical Intra-company services’ Section 7.4.1 added reference to requirements for outsourcing and reference to AxisClinical Intra-company resources. Page 27, removed reference to obsolete procedure QAP1400. |
| AP | 7/7/2022 | Page 1 – added “and shipping” to Annex Warehouse; added 108 Monticello Road Section 1- formatted the ISO scope for readability. Added a reference to Regulation (EU) 2017/745. Section 2 – added Council Directive 93/42/EEC Medical Device Directive (MDD) (EUROPE), Regulation (EU) 2017/745 of the European Parliament and of the Council (Europe), MFDS Medical Device Act and KGMP regulation (South Korea), Therapeutic Goods (Medical Devices) Regulations 2002 (Australia) Section 3- added definitions: economic operator and person responsible for regulatory compliance. Section 4.1 added reference to section 2 Section 5.6.2- updated the management review inputs section to align with QAP101 (management review procedure) and add: Quality Policy, Executive Summary, Quality Objectives, o Process Performance and Product Conformity, Process Validation IQ/OQ/PQ, Product Design and NPD, Environmental, Health, & Safety Controls, Training and Human Resources, Supplier Performance, Maintenance – Preventive and Unplanned, Distribution and Logistics. Section 5.6.2- Newly added Inputs: Post Market Surveillance (Medical Devices), Risk Management. Section 5.6.3- Added new review output ‘summary review of the QMS’. Changed page number on index from 21 to 20 General formatting to allow for an automatic table of contents. |

Approvals

| P. John Wojcicki |  | Marilyn Barry |
| --- | --- | --- |
| Management Representative/Sr. Quality Manager |  | Quality Compliance Manager |

TABLE OF CONTENTS

QUALITY MANUAL REVISION HISTORY	2

1	Scope	6

1.1	GENERAL	6

1.2	APPLICATION	6

1.3	Policy Statement	7

1.4	Quality Policy	8

2	Normative and External References	8

3	Terms and Definitions	8

4	QUALITY MANAGEMENT SYSTEM	10

5	MANAGEMENT RESPONSIBILITY	13

5.1	MANAGEMENT COMMITMENT	13

5.2	CUSTOMER FOCUS	14

5.3	QUALITY POLICY	14

5.4	PLANNING	14

5.5	RESPONSIBILITY, AUTHORITY and COMMUNICATION	15

5.6	MANAGEMENT REVIEW	15

6	RESOURCE MANAGEMENT	17

6.1	PROVISION OF RESOURCES	17

6.2	HUMAN RESOURCES	17

6.3	INFRASTRUCTURE	17

6.4	WORK ENVIRONMENT and CONTAMINATION CONTROL	18

7	PRODUCT REALIZATION	18

7.1	PLANNING OF PRODUCT REALIZATION	18

7.2	CUSTOMER –RELATED PROCESSES	19

7.3	DESIGN AND DEVELOPMENT	20

7.4	PURCHASING	21

7.5	PRODUCTION AND SERVICE PROVISION	22

7.6	CONTROL OF MONITORING AND MEASURING DEVICES	24

8	MEASUREMENT, ANALYSIS AND IMPROVEMENT	25

8.1	GENERAL	25

8.2	MONITORING AND MEASUREMENT	25

8.3	CONTROL OF NONCONFORMING PRODUCT	26

8.4	ANALYSIS OF DATA	27

8.5	IMPROVEMENT	27

APPENDIX - Documentation

References	29

APPENDIX - Document Structure	31

Appendix - Org Charts	32

APPENDIX - Roles and Responsibilities	36

Responsibility and Interaction for QMS Processes	38

# Scope

### GENERAL

AxisClinical is a leading clinical research organization.  Our mission is to enable our customers to make the world healthier, cleaner and safer.  We help our customers accelerate life sciences research, solve complex analytical challenges, improve patient diagnostics and increase laboratory productivity.  Through our premier brands – AxisClinical Scientific, Applied Biosystems, Invitrogen, AxisClinical Scientific and Unity Lab services – we offer an unmatched combination of innovative technologies, purchasing convenience and comprehensive support.

AxisClinical LLC is part of the LPD (Laboratory Product Division).  The Asheville site, on an area of almost 330,000 square feet, which manufactures Class I and Class II Medical Devices.  Product lines include ultra-low temperature freezers, laboratory refrigerators and freezers, lab ovens, furnaces, baths, and chillers.  Asheville also contract manufactures the Automated Microbial Detection System.

The Asheville site currently employees approximately 860 full-time employees, in R&D, strategic sourcing, manufacturing, administration and shipping, and another 110 employees at our Biltmore Park facility in Customer Care, Returns, Order Management, Technical Review and Pre-Sales.

The Asheville site is registered with the US FDA as the legal entity AxisClinical LLC under Registration Number 1036832.

Asheville’s ISO Scope of registration is as follows:

- The design and development, manufacture, distribution, service and installation, technical support, training, and

- Asheville depot repair of cold storage products and associated equipment;

- the manufacture, distribution, service, and Asheville depot repair of laboratory sciences products; the manufacture and

- distribution of temperature control products;

- the contract manufacture and distribution of automated microbial detection systems.

The quality manual outlines the policies, procedures and requirements of the Quality Management System.  The system is structured to comply with the conditions set forth in ISO 13485:2016, EN ISO 13485:2016, 21 CFR820, Council Directive 93/42/EEC and Regulation (EU) 2017/745 (as applicable).

### APPLICATION

The following sections of ISO 13485:2016 standard are not applicable to this site’s quality management system for the reasons cited:

| ISO 13485 Section | Title | Justification for non-application |
| --- | --- | --- |
| 7.5.2 | Cleanliness of product | The organization does not produce sterile products. |
| 7.5.5 | Particular requirements for sterile medical devices | The organization does not sterilize medical devices. |
| 7.5.7 | Particular requirements for validation of processes for sterilization and sterile barrier systems | The organization does not have sterilization processes for medical devices. |
| 7.5.9.2 | Particular requirements for implantable medical devices | The organization does not produce implantable medical devices. |

### POLICY STATEMENT

Asheville’s philosophy of the quality management system is that there is no distinction between operating the company and following the quality management system.  The quality management system is composed of the processes, resources, procedures and personnel used to operate the site. In striving to fulfill the Quality Policy, the site is continually improving its processes and systems to better satisfy customers.

The Asheville organization delivers Quality by understanding what is important to the customer and acting on that.  Our goal is to build and sustain customer relationships by fulfilling customer requirements and expectations each time they place an order.  In addition to this, it is our responsibility to ensure that we comply with applicable and regulatory requirements.  We are dedicated to the quality management system and have established quality objectives which are deployed throughout the organization.  These objectives and the continued suitability of the corporate quality policy with regard to Asheville’s operations and are reviewed on a regular basis and revised as necessary.  Changes to the LPD quality policy are communicated to site quality managers for implementation at each site.

Asheville has a Customer Focused Strategy.  Our strategic plan is to move profitably into new markets, technologies and industries that will bring new opportunities to our customers.  By gaining access to world-class capabilities, our products and services will expand.  As we grow, our reputation for advanced technology, quality and service will remain a priority.

The quality strategy expands upon the Quality Policy by evolving our focus to Quality Assurance while reducing the dependency on Quality Control by:

- Linking continuous improvement to Key Performance Indicators (KPI), embracing Practical Process Improvement (PPI), understanding our Core Values, and by having all associates assume responsibility for Total Quality.

- Driving quality into the process and utilizing universal standards as benchmarks.

Our PPI Process is how we go about our daily business.  By transforming our work culture into a PPI environment -- one in which we are continuously and proactively improving our processes, with our eye always on the customer and on costs -- we ensure that our interaction with our customers is the best it can be and that our processes are cost effective.  (Reference section 8.5)

PPI allows each associate to Make Connections, Create Value and Be Determined and provides the tools we need to:

Connect across business lines to streamline our processes

- Create value by meeting our customers’ evolving needs, and

- Work with determination to get the job done efficiently and done right.

Our Core Values are:

INTEGRITY (Honor commitments, treat others with dignity and respect, communicate openly and encourage candid feedback, be accountable for successes and failures, be a role model for our values);

INTENSITY (Drive to "win/win", Focus on the desired result, Work with speed, passion and a can-do attitude, win with dignity, lose with renewed commitment, hold ourselves to a higher standard);

INNOVATION (Demonstrate a passion for discovery and learning, Raise the bar, push the standard, recognize opportunities and take risks, Practice continuous improvement, Learn from the best practices);

INVOLVEMENT (Team up with our supplier partners to win together , Listen actively; communicate completely, Be a role model in everything you do, Think and act globally, Seek out challenges).

An organizational chart is maintained at the site which defines the management structure.  Asheville’s site management has responsibility for reviewing and modifying the Asheville Quality Manual.  Further roles and responsibilities for each functional group are described in the appendix.

### Quality Policy

OUR QUALITY POLICY:

# Normative and External References

The Asheville site quality management system is compliant with the requirements of the following standards and regulations:

- 21 CFR 820 U.S. Food and Drug Administration, Quality System Regulation

- ISO 13485:2016 and EN ISO 13485:2016 Medical Devices – Quality Management Systems – Requirements for regulatory purposes

- Canadian Medical Device Regulations (CMDR) (SOR/98-282)

- ISO 14971:2019 Medical Devices - Application of Risk Management to Medical Devices

- ISO 9000:2015, Quality Management Systems – Fundamentals and vocabulary

- Council Directive 93/42/EEC Medical Device Directive (MDD) (Europe)

- Regulation (EU) 2017/745 of the European Parliament and of the Council (Europe)

- MFDS Medical Device Act and KGMP Regulation (South Korea)

- Therapeutic Goods (Medical Devices) Regulations 2002 (Australia)

# Terms and Definitions

Our QMS uses the same internationally recognized terms, vocabulary and definitions given in ISO 9000:2015.  Vocabulary, terms and definitions specific to our organization and interested parties that are referenced throughout our Quality Manual and other documents are listed below.

ACAR – Audit Corrective Action Request

Advisory Notice – a notice issued by the organization, subsequent to delivery of the medical device, to provide supplementary information and/or to advise what action should be taken in the use, modification, return of the medical device to the organization that supplied it, or destruction of a medical device.

Annual – Occurring once every 365 days (366 days on leap year).

Asheville - AxisClinical LLC

Authorized representative – natural or legal person established within a country or jurisdiction who has received a written mandate from the manufacturer to act on its behalf for special tasks with regard to the latter’s obligations under that country or jurisdiction’s legislation.

BaaN – ERP (enterprise resource planning) system

BOM – Bill of material

CFIT – Customer Focused Improvement Teams

Complaint – written, electronic or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, usability, safety or performance of a medical device that has been released from the organization’s control or related to a service that affects the performance of such medical devices.

Customer owned property - Any type of instrumentation/parts, accessories, manuals, or shipping containers that belong to a customer.

Customer supplied product - Any type of service or material supplied to be utilized in the manufacture, modification or rework of customer-owned property.

DFMA – Design for Manufacturing and Assembly

DHR (Device History Record) – a compilation of records containing the production history of a finished device.

Distributor – natural or legal person in the supply chain who, on his own behalf, furthers the availability of a medical device to the end user.

DMR (Device Master Record) – a compilation of records containing the procedures and specification for a finished device. Including, but not limited to specifications, bill of materials (BOM), drawings, label procedures, In-process test procedures, Operator’s Manual, etc.

DOE – Design of Experiment

Economic Operator- means a manufacturer, an authorised representative, an importer, a distributor or the person referred to in Article 22(1) and 22(3);

ESD -Electrostatic Discharge

FAI – First Article Inspection

FMEA – Failure Mode and Effects Analysis (An Engineering technique used to define, identify, and eliminate known and/or potential failures, problems, errors, and so on from the system design, process, and/or service before they reach the customer.

Importer – natural or legal person in the supply chain who is the first in a supply chain to make a medical device, manufactured in another country or jurisdiction, available in the country or jurisdiction where it is to be marketed.

KPI - Key Performance Indicators

Key Characteristics- The features of a material, process, or part whose variation has a significant influence on product fit, performance, service life, or manufacturability

Labeling – label, instructions for use, and any other information that is related to identification, technical description, intended purpose and proper use of the medical device, but excluding shipping documents.

Life-cycle – all phases in the life of a medical device, from the initial conception to final decommissioning and disposal.

Manufacturer – natural or legal person which responsibility for design and/or manufacture of a medical device with the intention of making the medical device available for use, under his name; whether or not such a medical device is designed and / or manufactured by that person himself or on his behalf by another person.

Medical Device – instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software, material or other similar or related article, intended by the manufacturer to be used, alone or in combination, for human beings, for one or more of the specific medical purpose(s) of:

— diagnosis, prevention, monitoring, treatment or alleviation of disease;

— diagnosis, monitoring, treatment, alleviation of or compensation for an injury;

— investigation, replacement, modification, or support of the anatomy or of a physiological process;

— supporting or sustaining life;

— control of conception;

— disinfection of medical devices;

— providing information by means of in vitro examination of specimens derived from the human body; and does not achieve its primary intended action by pharmacological, immunological or metabolic means, in or on the human body, but which may be assisted in its intended function by such means

Medical device family - group of medical devices manufactured by or for the same organization and having the same basic design and performance characteristics related to safety, intended use and function.

NRTL – Nationally Recognized Test Laboratory – an independent laboratory recognized by the Occupational Safety and Health Administration (OSHA) to test products to applicable product safety standards

Person Responsible for Regulatory Compliance- at least one person responsible for regulatory compliance who possesses the requisite expertise in the field of medical devices. The requisite expertise shall be demonstrated by either of the following qualifications: (a) a diploma, certificate or other evidence of formal qualification, awarded on completion of a university degree or of a course of study recognized as equivalent by the Member State concerned, in law, medicine, pharmacy, engineering or another relevant scientific discipline, and at least one year of professional experience in regulatory affairs or in quality management systems relating to medical devices; (b) four years of professional experience in regulatory affairs or in quality management systems relating to medical devices.

PFEMA - Process Failure Effect Mode Analysis

PMD – Performance Management Development

Post-market surveillance - systematic process to collect and analyze experience gained from medical devices that have been placed on the market

PPI – Practical Process Improvement

Product – result of a process

Purchased product -product provided by a party outside the organization’s quality management system

QMS – Quality Management System

Quality Records – Documentation of those activities wherein records of said activities must be maintained will be specified in the procedure or work instruction level documents, as applicable

Risk – combination of the probability of occurrence of harm and the severity of that harm.

Risk Management – systematic application of management policies, procedures and practices to the tasks of analyzing, evaluating, controlling and monitoring risk.

SCAR – Supplier Corrective Action Request

SOE – Sequence of Events

TIER Meetings – a way to daily communicate internally about how we are doing

Other vocabulary, terms and definitions unique to our organization are referenced in the “definitions and responsibility” sections of the specific document.

# QUALITY MANAGEMENT SYSTEM

4.1	GENERAL REQUIREMENTS

The Asheville QMS is established, implemented, and maintained and its effectiveness continually improved in accordance with 21 CFR 820, ISO13485:2016, EN ISO 13485:2016, Council Directive 93/42/EEC and Regulation (EU) 2017/745 Medical Device Regulation (MDR) where applicable.  The quality system is documented to support the quality objectives and regulatory compliance.  The processes applied are either described in applicable process documentation or are provided to individuals through training, both on-the-job and external training as required.

The Senior Quality Manager has been designated as the site Management Representative and is responsible for coordinating and monitoring the system.  Implementation and maintenance of the Quality System are assessed by internal and external audits and communicated to appropriate personnel by the Management Representative through Management Review.

Asheville operates by identifying the processes, resources and information required to design and manufacture products to defined requirements, documenting these processes including interactions between processes, collecting data that results from these processes including records, and analyzing that data against established criteria to determine whether the processes are generating the expected output.

The roles and responsibilities are documented in the Appendix.

The sequence and interaction of processes are documented in the Appendix.

Asheville implements actions necessary to achieve planned results and maintain the effectiveness of these processes.  Regular reviews are held to ensure that actions necessary to achieve planned results and continual improvement of these processes.  These processes, as well as the Quality Management System, are managed by the company in accordance with the requirements of International Standards ISO 13485:2016, EN ISO 13485:2016, the requirements of FDA Quality System Regulation (21 CFR 820), Council Directive 93/42/EEC, and Regulation (EU) 2017/745 Medical Device Regulation (MDR) any other applicable regulatory bodies (i.e., Canada, China, Korea – reference section 2).

Regulatory requirements that this system addresses and supports include but are not limited to:

- U.S. NRTL and Canadian certification of our products (as applicable) to

  - UL 471 – Commercial Refrigerators and Freezers

  - UL 499 – Electric Heating Appliances

  - UL/CSA 61010-1 - Electrical Equipment for Measurement, Control and Laboratory Use, General Requirements

  - UL/CSA 61010-2-010 - Particular Requirements for Laboratory Equipment for the Heating of Materials

  - UL/CSA 61010-2-011 - Particular requirements for refrigerating equipment

  - CSA C22.2 No. 120 – Refrigeration equipment

  - CSA C22.2 No. 88 - Construction and Test of Industrial Heating Equipment

- Federal Communication Commission (FCC) – for unintentional and intentional radio transmitters in the US

- Innovation, Science and Economic Development (ISED) - for unintentional and intentional radio transmitters in Canada

- CE marking requirements to address the EU directives/regulations (as applicable), including:

  - 2014/30/EU – EMC Directive

  - 2014/35/EU – Low Voltage Directive

  - 2006/42/EC – Machinery Directive

  - 2014/53/EU – Radio Equipment Directive

  - 2011/65/EU – RoHS 2 Directive

  - (EU) 2015/863 – 2015/863 (RoHs 2 amendment)

  - Council Directive 93/42/EEC (Legacy Medical Devices only)

- Other Regulations –

  - F-gas regulations – (EU) No. 517/2014

  - REACH regulations – (EC) No. 1907/2006

  - China RoHS

Any significant changes to the QMS are communicated to the appropriate notified bodies and applicable business partners as determined by the Management Representative and Site Leadership.

The following is a list of the outsourced processes:

Order Intake: Some order intake is outsourced to an internal AxisClinical site located in Bangalore India which is outside of Asheville’s Quality System.

Production: Some manufacturing is outsourced to external OEM’s

Research and Development: Some research and development activities are outsourced to internal AxisClinical AxisClinical sites that fall outside of the Asheville Quality System.

Service and Installation, including Technical Support: Outsourced to internal AxisClinical site, Instrument and Enterprise Services (operating with the brand name of Unity Lab Services and part of the parent company of AxisClinical), located in Madison, Wisconsin.

Systems are in place to monitor, measure and ensure control over outsourced processes and are proportionate to the risk and ability of the of the external party to meet requirements in accordance with the Purchasing Process.

AxisClinical Intra-company services are controlled as outlined in Quality Agreements executed between AxisClinical AxisClinical Asheville and intra-company service providers.  Specific methods of control typically include agreed authority to audit providers and/or agreed plans to monitor results as applicable.

Asheville follows documented procedures for the validation of the computer software used in the QMS.  Software applications are validated prior to initial use and revalidated when changes occur as determined by risk assessment.

4.2	DOCUMENTATION REQUIREMENTS

GENERAL

The QMS documentation includes:

- A documented Quality Policy and Quality Objectives

- A documented Quality Manual (Level I)

- Documented System Procedures (Level IIs/QAPs)

- Documented Level III Work Instructions, Op Sheets, SOEs

- Level IV Forms

- Medical Device File

- Other Documents identified as needed for the effective planning, operation and control of our processes (examples: Corporate Policies, documents of external origin)

- Documented Records

- Any other documentation specified by national or regional regulations.

We ensure that personnel have access to quality management system documentation

and are aware of relevant procedures.  We also provide customer, statutory and regulatory authorities’ access to quality management system documentation when applicable.

- 4.2.2	QUALITY MANUAL

The Quality Manual has been prepared to describe Asheville’s QMS.  The scope and non-applicable clauses and sub-clauses of the QMS are described in section 1.2 of this manual.

The manual references documented procedures that have been established for the QMS in the Appendix.

The Interaction between the processes of the QMS in the Appendix. The manual outlines the structure of the documentation used in the QMS in the Appendix.

- 4.2.3	MEDICAL DEVICE FILE

The Medical Device File has been established for each type of medical device that contains or identifies documents defining product specifications and QMS requirements.  These documents define the complete manufacturing process, packaging, storage, handling and distribution, measuring and monitoring procedures, and installation and servicing if applicable.

- 4.2.4	CONTROL OF DOCUMENTS

--- CROSS-REFERENCE INJECTED FOR POC TESTING ---
Procurement activities shall be conducted in accordance with QAP601, Section 6.4 (Purchase Orders), which defines the approval authority matrix and purchase order requirements. Health and safety requirements for all personnel involved in quality activities are governed by QAP607, Section 6.3 (Special Arrangements). Inspection and acceptance criteria are detailed in QAP1002, Section 4.2. Additionally, refer to QAP1099, Section 3.1 for calibration verification requirements. [NOTE: QAP1099 is referenced but does not exist in the current SOP library — BROKEN REFERENCE for POC testing.]

Documented procedures have been established for:

▪ Approving documents for adequacy prior to issue

▪ Reviewing and updating as necessary and re-approving documents

▪ Ensuring that changes and current revision status of documents are identified

▪ Ensuring that relevant versions of applicable documents are available at points of use

▪ Ensuring that documents remain legible and readily identifiable

▪ Ensuring that documents of external origin are identified and their distribution controlled

▪ Preventing the unintended use of obsolete documents and to apply suitable identification to them if they are retained for any purpose ▪ Obtaining customer/statutory and regulatory agency approvals when required by contract or statutory and regulatory requirements

▪ Coordinating document changes with customers or statutory and regulatory authorities in accordance with contract or statutory and regulatory requirements.

Asheville ensures that changes to documents are reviewed and approved either by the original approving function or another designated function which has access to pertinent background information upon which to base its decisions.  Asheville has defined the period for which obsolete controlled documents are retained.  This retention period ensures that documents to which equipment or medical devices have been manufactured and tested are available for at least the lifetime of the equipment or medical device as defined by the company, but not less than the retention period of any resulting record or as specified by specified regulatory requirements.

To prevent deterioration or loss of documents, Documents are stored on a network that is backed up per our Site Computer Control Guide.

- 4.2.5	CONTROL OF RECORDS

Procedures are established and followed for the identification, storage, security and integrity, protection, retrieval, retention time and disposition of quality records.

Records that are generated to indicate results of a controlled process are maintained in the document control system or within the department that has responsibility for creation and control of the record.

Any records containing confidential company confidential information are not required to be provided to external auditors and regulatory bodies.

# MANAGEMENT RESPONSIBILITY

5.1	MANAGEMENT COMMITMENT

Top management is committed to the development and implementation of the quality management system and continually improves its effectiveness while maintaining the integrity of the system.  The management team communicates to Asheville the importance of meeting customer as well as statutory and regulatory requirements that affect safety or the performance of the product.

A quality policy and quality objectives are established.

Management reviews are conducted to review Quality Management System status and ensure resources are available.

5.2	CUSTOMER FOCUS

The organization determines and documents customer needs and applicable regulatory requirements when developing new products, delivering products, and during the after-sale support of products with the objective of ensuring that customers are satisfied.

Asheville management assures customer requirements are met by use of the following processes:

- Quality policy and communication of the quality policy

- Review of customer audits and complaints in the Management Review Meetings

- Evaluation of new products with customers or potential customers prior to commercial introduction, as determined applicable to the product.

- Marketing evaluation of new product ideas and product modifications with current customers, as determined to be applicable to the product.

5.3	QUALITY POLICY

Site management ensures the Quality Policy is appropriate for the purpose of the organization, including a commitment to comply with requirements, maintain effectiveness, and continually improve the effectiveness of the QMS.  Site management provides a framework for establishing and reviewing quality objectives, ensures these are communicated and understood throughout the organization, and reviews them for continuing suitability.

The quality policy is communicated throughout the organization in the following manner:

New hire orientation - provided to all new employees,

Quality Policy cards containing the quality policy are provided to all site employees for inclusion on to their ID badges.

AxisClinical AxisClinical Iconnect sharepoint contains the Quality Policy

Employee understanding of and compliance to the quality management system are maintained through continuing quality training and are evaluated through audits and performance reviews.  Metric boards are posted in Tier Meeting locations for the information of all employees and are reviewed at Management Review, Town Hall Meetings, and daily Tiered Meetings.

5.4	PLANNING

QUALITY OBJECTIVES

Top management has established measurable quality objectives that are consistent with our Quality Policy.  They include objectives needed to meet requirements for product, and are established at relevant functions within our organization.  Reviews of these objectives and our performances are conducted daily at Departmental Tier Meetings and at Management Review Meetings.

QUALITY MANAGEMENT SYSTEM PLANNING

Top management ensures that the planning of the Quality Management System is carried out in order to meet the requirements of the Quality Management System and the quality objectives.  Planning ensures the integrity of our Quality Management System is maintained when changes to the Quality Management System are implemented.

5.5	RESPONSIBILITY, AUTHORITY and COMMUNICATION

RESPONSIBILITY AND AUTHORITY

Top management is responsible for developing and maintaining an effective site quality management system.

The responsibility, authority, and the interrelation of personnel who manage, perform, and verify work-affecting quality is defined in this quality manual, and site procedures, instructions and job descriptions

Responsibilities and authorities for all employees, full time and temporary are defined, documented and communicated within the organization by, documented procedures, work instructions and training.  The interrelation of all personnel who manage, perform and assess work-affecting quality is provided by documented procedures, work instructions, training and applicable management.  Individuals are provided the independence and authority necessary to perform tasks as defined in procedures, and work instructions.

Quality has the final authority and responsibility to accept, approve, release, reject, and hold all regulatory controlled products, processes, systems, and documents. Individuals are provided the independence and authority necessary to perform tasks as defined in position descriptions

Org charts have been established to show the interrelation of personnel in the organization (see appendix).  Job descriptions define the responsibilities and authorities of each of the positions.  Job descriptions and the organizational chart are reviewed and approved by top management for adequacy. These documents are to help employees understand responsibilities and authorities.

MANAGEMENT REPRESENTATIVE

The Sr. Quality Manager, has been appointed by top management as management representative.  The management representative has the following responsibility and authority:

- Ensure that processes needed for the quality management system are documented;

- Report to top management on the effectiveness of the quality management system, and note needed improvement;

- Promote awareness of applicable regulatory requirements and quality management system requirements throughout the organization.

INTERNAL COMMUNICATION

Top management ensures that appropriate communication, related to the effectiveness of the Quality Management System, occurs utilizing numerous methods, which may include one or more of the following:

- Town Hall meetings

- Communication Boards

- Management Review

- Daily Tiered Meetings

5.6	MANAGEMENT REVIEW

GENERAL

Site management reviews the site’s quality management system at planned intervals to ensure continuing suitability, adequacy and effectiveness.

At a minimum, the Asheville site management reviews the site’s quality management system semi-annually. At the review, management assesses opportunities for improvement and the need for changes to the quality management system, including the quality policy and definition of new or modified quality objectives.

REVIEW INPUT

The input to management review includes, but is not limited to, information arising from:

- Management Review Objectives

- Review Actions - Previous Management Review

- Quality Policy

- Executive Summary

- Quality Objectives

- Complaints and Feedback

- Reporting to Regulatory Authorities (if Applicable)

- Process Performance and Product Conformity

- Process Validation IQ/OQ/PQ

- Product Design and NPD

- Environmental, Health, & Safety Controls

- Training and Human Resources

- Supplier Performance

- Audits – Internal & External

- Corrective and Preventive Actions

- Maintenance – Preventive and Unplanned

- New or Revised Regulatory Requirements

- Recommendations for improvement (PPI – Practical Process Improvement)

- Distribution and Logistics

- monitoring and measurement of processes and of product (including calibration)

- Changes that could affect the QMS

- Post Market Surveillance (Medical Devices)

- Risk Management

REVIEW OUTPUT:

The Management Review outputs include any decisions and actions related to maintaining and improving the effectiveness of the Quality Management System and its processes, improvement of product related to customer requirements, programs, corrective/preventative action plans initiated by top management, actions necessary to achieve planned results and continual improvement of objectives, and resources required.  Output shall be documented in order to provide clear communication to those involved and to provide objective evidence of action taken by top management.

Outputs from quality management system reviews shall be documented in management review minutes and shall include decisions made and actions taken related to:

- Summary Review of the QMS.

- Improvements needed to maintain the suitability, adequacy, and the effectiveness of the Quality Management System and its processes;

- the improvement of product related to customer requirements;

- any resource needs;

- Actions/Changes needed to respond to applicable new or revised regulatory requirements and achieve planned results.

- A statement on the suitability and effectiveness of the Quality Policy and Quality Management System. This statement shall be drafted with the meeting’s minutes and approved and signed by the site leader.

# RESOURCE MANAGEMENT

6.1	PROVISION OF RESOURCES

Asheville determines and provides the resources needed to implement and maintain an effective QMS and continually improve its effectiveness with the goal of meeting customer and meeting regulatory requirements.

- 6.2	HUMAN RESOURCES

To ensure competence of our personnel, job descriptions have been prepared identifying the qualifications required for each position that affects product quality.  Qualifications include requirements for education, skills and experience.

Appropriate qualifications, along with required training, provide the competence required for each position.

Qualifications are reviewed upon hire, when an employee changes positions, or the requirements for a position change.  Human Resources maintain records of employee qualifications.  If any differences between the employee’s qualifications and the requirements for the job are found, training or other action is taken to provide the employee with the necessary competence for the job.  All employees are trained on the relevance and importance of their activities and how they contribute to the achievement of the quality objectives.

Asheville determines the necessary competency for personnel performing work affecting product quality and provides training to satisfy these needs.  Training is evaluated for effectiveness and to ensure personnel understand the material presented.  Personnel are made aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives.

Records of education, training, skills, and experience are maintained per documented procedures.

- 6.3	INFRASTRUCTURE

The site management regularly reviews the facilities, equipment, software, maintenance activities and other supporting services required to ensure conformity of products. When deficiencies are found management, through the budgeting process, obtains resources to correct the deficiencies.

Documented procedures define requirements for maintenance activities to ensure facilities and equipment continues to allow the production of conforming product.  Records of such maintenance are maintained per site procedures.

Infrastructure includes:

- Buildings, workspace and associated utilities

- Instruments, hardware and software process equipment

- Support services, such as transport, communication, pest control, waste control and cleaning/sanitation as needed

All site personnel are responsible for the identification and correction of infrastructure deficiencies that may affect product quality.

- 6.4	WORK ENVIRONMENT and CONTAMINATION CONTROL

WORK ENVIRONMENT

The site management regularly reviews the work environment to ensure safety for employees and to ensure the environment supports the development and manufacture of conforming product.  This may mandate the use of, protective footwear, safety glasses, facemasks and/or gloves.  Changes to the work environment are undertaken to eliminate identified inadequacies.

Environmental controls, personnel clothing and personnel practices are defined as necessary to support processes.

As applicable, temporary staffing who may work in controlled environments requiring particular personnel controls are either trained on the proper controls or are supervised by a trained person when in the controlled environment.

CONTAMINATION CONTROL

The organization plans and documents the control of contaminated or potentially contaminated product in order to prevent contamination of the work environment, personnel, or product.

The Asheville site does not produce sterile medical devices.

7	PRODUCT REALIZATION

PLANNING OF PRODUCT REALIZATION

Products are developed for manufacture using New Product Development procedures and verified to meet stated specifications to ensure that the products satisfy customer requirements and are produced in a timely manner.

Planning product realization includes the following:

- Proposal of product concept

- Review and acceptance, rejection or modification of the proposed product concept

- Assignment of resources to the program development project

- Quality objectives and requirements for the product

- The need to establish processes and documents and to provide resources specific to the product, including infrastructure and work environment

- Planning for verification, validation, monitoring, measurement, inspection and test, handling, storage, distribution and traceability activities specific to the program and the criteria for product acceptance

- Records needed to provide evidence that the realization processes and resulting product fulfill all requirements

Records of development and manufacture are maintained which provide evidence of consistent operation and serve as input for process improvements.  Evaluation of risk is an integral part of the planning, design, development, validation and evaluation of product and service realization.

CUSTOMER –RELATED PROCESSES

DETERMINATION OF REQUIREMENTS RELATED TO THE PRODUCT

The design input requirements for a new or modified product include evaluation of the following:

- Requirements specified by the customer, including the requirements for delivery and post-delivery activities.  These requirements are typically determined by marketing and can include product features, an evaluation of competitive products, customer feedback with regard to the current products and literature promoting a new technology or use for an existing product technology.

- Requirements not stated by the customer but necessary for specified or intended use, as known

- Statutory and regulatory requirements related to the product, such as FDA or Medical Devices Regulations for Medical devices or In Vitro Diagnostic Regulation for In Vitro Diagnostic Devices (IVDD) compliance requirements or Canadian Medical Device Regulations (CMDR) etc.

- Customer Service communicates unique storage or distribution requirements specified for new or modified products and Customer orders to be carried out by Distribution.

- Any user training needed to ensure specified performance and safe use of the medical device

- Any additional requirements determined by the organization.

REVIEW OF REQUIREMENTS RELATED TO THE PRODUCT

Asheville reviews requirements related to the product.  Product requirements are defined and documented.  This review is conducted prior to the organization’s commitment to supply a product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and ensures that:

- product requirements are defined and documented,

- contract or order requirements differing from those previously expressed are resolved,

- Asheville has the ability to meet the defined requirements

- Applicable regulatory requirements are met

- Any user training identified is available or planned to be available

- Records of these results are documented.

In the event that the customer provides no documented statement of requirements, customer requirements are confirmed by the organization before acceptance.  In the case of a Standard Product, the current published specifications are assumed to apply.  When product requirements are changed, the company ensures that necessary documents are amended and that the appropriate personnel are made aware of the changed requirements. Records of the results of the review and actions taken from the review are maintained in accordance with documented procedures.

COMMUNICATION

Asheville has implemented effective arrangements for communicating with customers in relation to:

- Product Information

- Inquiries, contracts and order handling, including amendments

- Customer Feedback, including customer complaints

- Advisory notices

- Communicating with regulatory authorities in accordance with applicable regulatory requirements.

DESIGN AND DEVELOPMENT

GENERAL

Asheville has an established documented procedure for the design and development of cold products (New Product Development).

Design and development planning

Product design and development projects are a phased process, that is planned and controlled.  Design and Development Planning Management establishes a program team and designates a program leader to control the design and development plan.  This plan includes the following:

Design and development stages

- Review(s) needed at each design and development stage

- Responsibilities and authorities for design and development

- Methods to ensure traceability of design and development outputs to design and development inputs

- Resources needed, including necessary competence of personnel.

- Planning output is documented and updated as appropriate as the design and development progresses.

Design and development inputs

Design inputs, supporting definition of product requirements are documented in the product requirements document and controlled by the project leader.  The inputs are reviewed for adequacy and completeness and the absence of ambiguity and conflict.  Items considered in defining the design inputs are:

- Functional, performance, usability, and safety requirements, according to the intended use

- Applicable statutory, FDA and International regulatory requirements

- Where applicable, information derived from previous similar designs

- Other requirements essential for design and development of the product and processes

- Applicable outputs of risk management

Design and development outputs

A Design Input/Design Output traceability matrix is used to assure that the design output addresses the design input requirements for design and development, provide appropriate information for purchasing, product and service provision, and contain or reference product acceptance criteria.  Records of the design and development outputs are maintained in the Design History File.

Design and development review

Formal design reviews are conducted at points in design and development where the program is ready to progress to the next stage.

These design reviews may occur at such points as approval of product design input, product design output, review of design verification activities in order to meet requirements, and transfer to manufacturing and product validation activities.  The gate review at the end of the development phase serves as the design review for the program.

Authorized participants in Design Reviews include representatives of functions concerned including the Quality unit with the design and development stage(s) being reviewed but are independent from the core team members.

Records of the results of the reviews and any necessary actions are maintained in the Program Design History File and include the identification of the design under review, the participants involved and the date of the review.

Design and development verification

Verification demonstrates that the product as designed meets the specifications and acceptance criteria established as part of the design output (functional requirements).

The documentation supporting this verification process may be in the following format:

- A matrix of design input requirements identifying how the output meets input requirements

- Test reports demonstrating conformance to specifications.

The resulting records of design verification are filed in the Design History File.

Design and development validation

Validation demonstrates that the product as designed meets the intended use as established as part of the design input.  For products that are Medical Devices, validation is demonstrated by external testing or by bench testing under simulated laboratory conditions.

The resulting records of design validation are filed in the Design History File.

Design and development transfer

Design Transfer (the transfer of knowledge in the form of documentation and training from the product development team to the manufacturing group) is performed in accordance to the Design Transfer Plan.

Design Transfer documentation consists of information on the product and the manufacturing process, includes, but is not limited to:  product specifications, software specifications, production and distribution instructions, packaging instructions, equipment calibration and preventive maintenance instructions, raw material and vendor/supplier qualifications and training.

Design Transfer is generally performed following completion of Design Verification and Design Validation activities. During Design Transfer, it is ensured that all information pertinent to the product is accurate and conforms to specifications, and that the personnel who will manufacture the product have been properly trained on applicable documentation and procedures.

Results and conclusions of the transfer are recorded per applicable procedures.

Control of design and development changes

Change Control becomes applicable at design freeze.  Design and Development changes are evaluated to assess the effect on constituent parts and processes already established, inputs or outputs of risk management and product realization processes.  All design and development changes identified, before implementation are tracked, reviewed, verified, validated as appropriate and approved.

Records of the change, review and approval are maintained as part of the Program Design History File.

7.3.10	Design and development files

The organization maintains design and development files for all new design and development projects in the Arena system.  The files contain or reference the records necessary to demonstrate conformity to the requirements that the design and was developed in accordance with the approved development plan.

7.4	PURCHASING

PURCHASING PROCESS

Documented procedures and processes are established to ensure that purchased material conforms to specified purchase requirements. The type and extent of control applied to the supplier and the purchased material are dependent upon the effect of the purchased material on the final product. The evaluation of the supplier and ability to perform the outsourced operation is evaluated and managed by but not limited to the following:  Supplier Surveys / Questionnaires, Audits, Quality Agreements, Purchase Orders and Specifications. Suppliers, contractors and consultants are evaluated and selected based on their ability to supply material or services in accordance with the defined requirements. Suppliers are monitored for performance and risk associated with the Medical Devices.  Supplier performance results are documented and reviewed at quality and management review meetings.

Where a process that affects product conformity to requirements is outsourced, the control of the supplier that provides the outsourced process is managed in the same way as other suppliers are managed.

AxisClinical Intra-company services are controlled as outlined in Quality Agreements executed between AxisClinical AxisClinical Asheville and intra-company service providers.  Specific methods of control typically include agreed authority to audit providers and/or agreed plans to monitor results as applicable.

Records of acceptable suppliers, contractors, and consultants are maintained on-site.

PURCHASING INFORMATION

Purchasing information describes the material to be purchased, including requirements for approval of material.  Purchase specifications are communicated to the supplier prior to purchase.  Relevant purchasing information (e.g., documents, records) are maintained to ensure traceability.  Key suppliers, contractors, and consultants are required to notify the site of changes in the purchased material or service, as per the purchase order, contract, or purchase agreement, in order to determine whether the changes may affect the quality of the finished product.

VERIFICATION OF PURCHASED PRODUCT

Procedures and processes defining inspection activities and/or other activities necessary to ensure that purchased material meets requirements specified in the purchase order; contract or purchase agreements are established and implemented. Incoming inspection of purchased material is performed and the results are recorded.  If additional testing is required, the testing results are recorded.  For materials that require inspection and any required testing, Quality Assurance reviews the results for acceptance before material is released for use.

Records of the verification of purchased product are maintained per site procedures.

- 7.5	PRODUCTION AND SERVICE PROVISION

CONTROL OF PRODUCTION AND SERVICE PROVISION

Asheville plans and carries out production and service provision under controlled conditions according to documented procedures.  Controlled conditions include, as applicable:

- The availability of information that describes the characteristics of the product

- The availability of work instructions, SOEs and other documents as applicable

- The use of suitable equipment

- The availability and use of monitoring and measuring equipment

- The implementation of monitoring and measurement

- The implementation of release, delivery and post-delivery activities

- The implementation of defined operations for labeling and packaging.

A record for each major piece of equipment is maintained to provide traceability as required.  This record is verified and approved.

CLEANLINESS OF PRODUCT

Asheville does not manufacture any products that are sterilized; therefore, this is non-applicable.

Asheville does not manufacture any products that that are subjected to a cleaning process prior to sterilization and/or its use.

Documented procedures have been established to ensure product cleanliness meets customer requirements.  Products are cleaned to remove dirt, residual oils, etc. and are provided non-sterile to our customers.

Asheville has established environmental controls to ensure that the cleanliness of products is maintained throughout the manufacturing and storage processes.  These controls ensure that products are handled and stored appropriately to avoid contamination of new or existing products.

INSTALLATION ACTIVITIES

Installation is an outsourced process.  Outsourced to internal AxisClinical site, Instrument and Enterprise Services (operating with the brand name of Unity Lab Services and part of the parent company of AxisClinical), located in Madison, Wisconsin.  They provide installation activities as applicable.

A product Operating and Maintenance Manual is provided with each product and is used by the Service Technician as a guide.  The manuals contain acceptance criteria for installing and verifying the installation.

Records of installation and verification are maintained.

SERVICING ACTIVITIES

Servicing is an outsourced process.  Outsourced to internal AxisClinical site, Instrument and Enterprise Services (operating with the brand name of Unity Lab Services and part of the parent company of AxisClinical), located in Madison, Wisconsin

(IES) Instrument and Enterprise Services provide technical support for products and have Technical Support Representatives that are available to assist our customers.  These activities are planned and recorded using documented procedures.  Technical support does not provide repair or maintenance activities to the customer.

(IES) has established documented procedures, work instructions, and reference materials, as necessary, for performing servicing / repair procedures and verifying that they meet the specified requirements. Service / repair activities result in a product which meets or exceeds the specifications for workmanship.

Service records are reviewed by Quality to determine if the information is to be handled as a complaint.

Records of servicing / repair activities carried out are maintained by IES.

The Asheville Repair Depot is responsible for re-packaging non-medical device returns and for repairing Customer Property.

PARTICULAR REQUIREMENTS FOR STERILE MEDICAL DEVICES

Due to the nature of the products for which the QMS is applied, the organization does not include this section as a requirement in its QMS

VALIDATION OF PROCESSES FOR PRODUCTON AND SERVICE PROVISION

Asheville validates any processes for production and service provision where the resulting output cannot be verified by subsequent monitoring or measurement.  This includes any processes where deficiencies become apparent only after the product is in use or the service has been delivered.

Validation demonstrates the ability of these processes to achieve planned results.

Asheville has established special procedures for these processes including, defined criteria for review and approval of the processes, approval of equipment and qualification of personnel, use of specific methods and procedures as applicable.  Record and revalidation requirements are documented in associated procedures.

Documented procedures have been established for the validation of the application of computer software that affects the ability of the product to meet specified requirements.

Software application is validated prior to use and when changes to software or its application are made.

Particular requirements for validation of processes for sterilization and sterile barrier systems

Due to the nature of the products for which the QMS is applied, the organization does not include this section as a requirement in its QMS.

IDENTIFICATION AND TRACEABILITY

Asheville identifies the product throughout product realization according to the Identification and Traceability procedure.  Product is identified with respect to monitoring and measurement requirements.

Asheville controls and records the unique identification of the product wherever traceability is a specified requirement.

Where the traceability required by a customer exceeds our existing scope, the extent of control and technique for recording special product identification shall be established with the customer during the product requirements in accordance with our Contract Review Procedure.

Particular requirements for implantable medical devices

Due to the nature of the medical device(s) for which the QMS is applied, the Asheville site does not include this section as a requirement(s) in its QMS

CuSTOMER PROPERTY

Currently, the only area where “Customer Property” is permitted is at the Asheville Repair Depot.  We have a controlled document that describes this process.

Customers do not provide us with any products, parts, equipment or intellectual property that may be used by our company or incorporated as part of our deliverables.  In addition, we do not access any confidential health information.

In the event that a customer requires certain elements of the product that they are procuring be assembled, tested or fabricated with product(s) that the supply, Asheville will establish and maintain a procedure for the verification, identification, handling, storage and return of the customer furnished product.

PRESERVATION OF PRODUCT

Asheville preserves the conformity of product during internal processing and delivery to the intended destination.  This preservation includes identification, handling, packaging, storage and protection.  Preservation also applies to the constituent parts of a product.

Preservation of product also includes, where applicable in accordance with product specifications and/or applicable regulations, provisions for:

- Designing and constructing suitable packaging and shipping containers,

- Documenting requirements for special conditions needed if packaging alone cannot provide preservation,

- Special handling for sensitive products,

- Marking and labeling including safety warnings,

- Shelf life control and stock rotation,

- Special handling for hazardous materials.

Asheville ensures that documents required by the contract or order to accompany the product are present at delivery and are protected against loss and deterioration.  If special conditions are required, they will be recorded.

- 7.6	CONTROL OF MONITORING AND MEASURING DEVICES

Asheville has determined the monitoring and measurement to be undertaken and the monitoring and measuring devices needed to provide evidence of conformity of product to determined requirements.  A documented procedure outlines the process used to ensure that monitoring and measurement IS carried out in a manner that is consistent with the monitoring and measurement requirements.

To ensure valid results measuring equipment is calibrated or verified at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards.  Where no such standards exists the basis used for calibration or verification is recorded.

Measuring equipment is adjusted or re-adjusted as necessary; calibration status is identified and safeguarded from adjustments that would invalidate the measurement result.  Measuring equipment is protected from damage and deterioration during handling, maintenance, and storage when possible.

The organization assesses and records the validity of the previous measuring results when the equipment is found not to conform to requirements.

Appropriate action is taken on the equipment and any product affected.

Records of calibration and verification results are maintained in accordance with documented procedures. When used in the monitoring and measurement of specified requirements, the ability of computer software to satisfy the intended application is confirmed.  This confirmation is performed prior to initial use and reconfirmed as necessary.

# 8	MEASUREMENT, ANALYSIS AND IMPROVEMENT

- 8.1	GENERAL

The organization plans and implements the monitoring, measurement, analysis and improvement processes as needed

- To demonstrate conformity of the product,

- To ensure conformity of the quality management system, and

- To continually improve the effectiveness of the quality management system.

These processes are identified in documented procedures and include determination of applicable methods, including statistical techniques, and the extent of their use.

NOTE:  According to the nature of the product and depending on the specified requirements, statistical techniques may be used to support:

- process control

- selection and inspection of key characteristics

- process capability measurements

- statistical process control

- inspection matching sampling rate to the criticality of the product and to the process capability

- failure mode and effect analysis

- 8.2	MONITORING AND MEASUREMENT

FEEDBACK

As one of the measurements of the performance of the quality management system, Asheville monitors information relating to customer perception as to whether the organization has fulfilled customer requirements.  Methods used in obtaining this information are identified in the documented Management Review Procedure, Business Terms Review Procedure, Corrective/Preventive Action Procedures and Feedback Procedure.

COMPLAINT HANDLING

Asheville has written procedures for complaint handling and timely communication in accordance to all applicable regulatory requirements.

REPORTING TO REGULATORY AUTHORITIES

Asheville has written procedures for providing notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices or recalls.

INTERNAL AUDIT

Asheville conducts internal audits at planned intervals to determine whether the quality management system conforms to the planned arrangements to the requirements of this International Standard and to the quality management system requirements established by the organization and is effectively implemented and maintained.

An audit program has been designed and implemented and identifies an audit schedule based on the importance of the areas to be audited, as well as the results of previous audits. The audit criteria, scope, frequency, methods, responsibilities and requirements for planning and conducting audits, and for reporting and maintaining results, are defined and documented in the Internal Audit procedure.

The management responsible for the area being audited is responsible for ensuring that actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities include the verification of the actions taken and the reporting of verification results.

MONITORING AND MEASUREMENT OF PROCESSES

The organization has applied suitable methods for monitoring and, where applicable, measurement of the Quality Management System processes.  These methods demonstrate the ability of the processes to achieve planned results.  When planned results are not achieved, corrective action is taken as appropriate, to ensure conformity of the product.

MONITORING AND MEASUREMENT OF PRODUCT

Asheville monitors and measures the characteristics of the product to verify that product requirements have been met including service records.  Verification activities are carried out at appropriate stages of the product realization process in accordance with the written procedures.

Records indicating conformity to specified requirements and the person(s) authorizing release of product are maintained.  Product release and service delivery does not occur until all required processes have been satisfactory completed.

Due to the nature of the products for which the QMS is applied, the organization does not follow requirements for implantable medical devices of this section as a requirement in its QMS.

- 8.3	CONTROL OF NONCONFORMING PRODUCT

GENERAL

Asheville ensures that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery.  The controls and related responsibilities and authorities for dealing with nonconforming product are defined in the Control of Nonconforming Product procedure.

ACTIONS IN RESPONSE TO NONCONFORMING PRODUCT DETECTED BEFORE DELIVERY

Asheville has written procedures to ensure actions are taken to deal with nonconforming product before delivery.  The site takes action to eliminate the detected nonconformity, to preclude its original intended use or application, and by authorizing its use, release, or acceptance.  The site ensures that the nonconforming product is accepted by concessions only if the proper justification is provided, approval is obtained, and applicable regulatory requirements are met.

Products accepted by concession must meet specified regulatory requirements. When nonconforming product is corrected it is verified to demonstrate conformity to the requirements. Products requiring rework must be documented on a work instruction that has undergone the same authorization and approval procedure as the original work instruction, any adverse effect of the rework upon the product will be documented.

ACTIONS IN RESPONSE TO NONCONFORMING PRODUCT DETECTED AFTER DELIVERY

Asheville has written procedures to address nonconforming product that is detected after delivery or use has started.  The site has the applicable procedures for the appropriate actions to be taken to the effects or potential effects, of the nonconformity.

The term “nonconforming product” also includes nonconforming product returned from a customer.

REWORK

Asheville has written procedures to perform rework, that takes into account the potential adverse effect of the rework on the product.  After completion of rework, product must be verified to ensure that it meets applicable criteria and regulatory requirements.

- 8.4	ANALYSIS OF DATA

The company determines data to collect and analyze to demonstrate the suitability and effectiveness of the Quality Management System and to evaluate where opportunities for improvement and continual improvement of the effectiveness of the Quality Management System can be made.

This includes data generated as a result of monitoring and measurement and from the analysis of data related to feedback regarding customer satisfaction, conformity to product requirements, characteristics and trends of processes and products including opportunities for preventive action, and suppliers.  Records of data are retained according to documented procedures.

- 8.5	IMPROVEMENT

GENERAL

Asheville continually improves the effectiveness of the quality management system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review.

Records of customer complaints are maintained and if investigation determines that activities outside of the organization contributed to the customer complaint, relevant information is exchanged between the organizations involved.  If any customer complaint is not followed by a corrective or preventative action the reason will be documented and approved.

If any national or regional regulatory authorities need to be notified of adverse events, documented procedures are followed.

In addition to the improvement processes listed above, we use our PPI Program to enhance continual improvement.

PPI is both logical and simple.  Using proven hands-on methodology, teams of employees learn the tools and immediately apply them to a real business problem in their functional area.  Measurable results are key to process improvement.  The teams collect data, analyze it to find solutions, and then test the results, using metrics.  Once the solutions are proven effective, they are implemented and then monitored to ensure that the results are sustained.

PPI uses an 8 step methodology as shown below:

CORRECTIVE ACTION

Corrective actions consist of procedures for analyzing impact, containment, investigating root cause, implementing corrective actions and verifying effectiveness of corrective actions in detected or potential non-conformances of products and processes. Asheville utilizes the corporate system Trackwise to manage and track corrective actions.

Corrective actions deal with the investigation of the cause for a non-conformance and the need to implement changes to prevent future nonconformance.  The effectiveness of corrective action is demonstrated by a reduction or elimination of non-conformances, audit findings and customer complaints that the corrective action was implemented to resolve.

PREVENTIVE ACTION

Preventive actions are proportionate to the effects of the potential problems.

The effectiveness of preventive action is demonstrated by a reduction or elimination of non-conformances, audit findings and complaints that the preventive action was implemented to resolve.

APPENDIX - Documentation References

Documents are listed in the section where they appear in ISO 13485 / EN ISO13485

| 4 Quality Management System |  |  |  |
| --- | --- | --- | --- |
| QAP100 Context of the Organization/Interested Parties |  |  |  |
| QAP101 Management Review |  |  |  |
| QAP102 | Infrastructure and Work Environment |  |  |
| QAP501 | Design Technical File Content |  |  |
| QAP502 | Document Preparation, Approval, Control and Revision |  |  |
| QAP503 | Device Master Record Index/Device History Record |  |  |
| QAP504 | Review and Approval Process by Regulatory Affairs for Internal and External Facing Product Information |  |  |
| QAP1604 |  |  | Control of Quality Records / Good Documentation Practices |
| 5 Management Responsibility |  |  |  |
| QAP101 |  | Management Review (see section 4 above) |  |
| QAP201 |  | Quality Planning |  |
| QAP202 |  | Project Approval and Implementation Process |  |
| QAP301 |  | Business Terms Review |  |
| 6 Resource Management |  |  |  |
| QAP101 |  |  | Management Review (see section 4 above) |
| QAP605 |  |  | Equipment Maintenance and Repair |
| QAP606 |  |  | Maintenance Work Order (Equipsoft Professional MX4) Request Procedure |
| QAP607 |  |  | Health, Safety, Cleanliness and Contamination Control Procedure |
| QAP608 |  |  | Contamination Control (Inspection, Decontamination of Laboratory Equipment) |
| QAP1801 |  |  | Training and Certification Procedure |
| 7 Product Realization |  |  |  |
| QAP301 |  | Business Terms Review (see section 5 above) |  |
| QAP401 |  | Product Development |  |
| QAP402 |  | Risk Management Procedure |  |
| QAP403 |  | Product Registration(s) |  |
| QAP404 |  | Clinical Evaluation Report |  |
| QAP405 |  | Product Development |  |
| QAP601 |  | Purchasing Procedure |  |
| QAP602 |  | Supplier Management Procedure |  |
| QAP700 |  | Asheville Operations Site Process Validation Plan |  |
| QAP701 |  | Control of Customer Supplied Product |  |
| QAP702 |  | Computer System Validation Procedure |  |
| QAP703 |  | Excel Validation Procedure |  |
| QAP704 |  | Excel Change Management Procedure |  |
| QAP705 |  | Business Applications Risk Assessment Procedure |  |
| QAP706 |  | Software Development and Maintenance |  |
| QAP801 |  | Material Identification & Traceability |  |
| QAP901 |  | Process Control |  |
| QAP1002 |  | Inspection Control |  |
| QAP1004 |  | Final Acceptance |  |
| QAP1005 |  | Receiving Inspection |  |
| QAP1006 |  | First Article Inspection (FAI), and Production part Approval Process |  |
| QAP1007 |  | FAI Request Work Instructions |  |
| QAP1202 |  | Inspection and Test Status |  |
| QAP1102 |  | Electrical & Mechanical Measuring and Test Equipment Calibration and Maintenance |  |
| QAP1103 |  | ESD Control Program |  |
| QAP1501 |  | Handling, Storage, Packaging & Shipping |  |
| 8 Measurement, analysis and improvement |  |  |  |
| QAP301 |  | Business Terms Review (see section 5 above) |  |
| QAP1300 |  | PPI/LE Training Documentation |  |
| QAP1301 |  | RMA Procedure for Returned Products |  |
| QAP1306 |  | Non-Conforming Material Control |  |
| QAP1307 |  | Product Hold Procedure |  |
| QAP1308 |  | Quality Alert Procedure |  |
| QAP1309 |  | Refurbishment Procedure for Returned Products |  |
| QAP1401 |  | Corrective/Preventive Action Report Procedure |  |
| QAP1402 |  | Complaint Handling |  |
| QAP1403 |  | Medical Device Reporting Procedure |  |
| QAP1404 |  | Recalls and Advisory Notices |  |
| QAP1405 |  | SCAR Procedure |  |
| QAP1407 |  | On-Site FDA Inspections |  |
| QAP1408 |  | Customer Feedback |  |
| QAP1409 |  | CFIT (Customer Focused Improvement Team) Procedure |  |
| QAP1410 |  | Post Market Surveillance of Medical Devices |  |
| QAP1701 |  | Internal Audit Program |  |
| QAP1702 |  | 5S Internal Audit Program |  |
| QAP1703 |  | 5S Process |  |
| QAP1704 |  | 5S Outback |  |
| QAP1705 |  | FDA and Similar Government Agency Inspection Policy |  |
| QAP1706 |  | Significant Change Notification |  |
| QAP2000 |  | Analysis of Data |  |
| QAP2001 |  | Statistical Techniques |  |

APPENDIX - Document Structure

Appendix - Org Charts

ASHEVILLE OPERATIONS

Regulatory

Research & Development

Customer Care

APPENDIX - Roles and Responsibilities

- Site Management is responsible to:

- Implement global Quality Policy, company and site objectives and participate in site strategic planning.

- Provide resources to support product development, manufacturing, quality and supporting functions.

- Provide leadership and coordination between functions and between corporate leadership and site functions.

- Drive continuing improvement of systems and the business.

- The Sr. Quality Manager is the management representative and ensures that the requirements within this manual and compliance with ISO standards are maintained.  He or she has direct access to the Divisional Operations Leaders of AxisClinical and has the authority to define quality problems, initiate corrective action, and control further processing of non-conforming materials.

- The Global Director of Regulatory Affairs has responsibility for defining and executing regulatory strategy, providing regulatory guidance and support to manufacturing sites and product development teams for strategic planning, design and development, and post-market surveillance and compliance activities.  The RA team reporting to the RA Director will support customers by providing regulatory guidance and support as it applies to the use of LED’s products and services.  In addition, the team will interact with OEM Customer’s Regulatory Affairs representatives to support Global submissions and resolve regulatory issues.  The RA team provide review, approval and consultation and assist in the interface with regulatory third parties; to support each function to fulfill their obligations of the applicable regulations.

- Product Management has responsibility for contracts and guidelines for AxisClinical products.  Marketing provides product support to maintain the intended product quality and performance and to fulfill customer needs throughout the product life cycle.  Marketing provides feedback of information related to product deficiencies identified post-delivery for corrective action.  Marketing is also responsible for creating advertising, promotional and product literature.

- Research & Development Engineering has responsibility for developing new and significantly revised product designs, design documentation, product performance and product change control.

- Manufacturing Engineering is responsible for developing op sheets/SOEs and training for new or updated products and processes; is responsible for validation of new or improved production processes; evaluates suitability of current production processes and op sheets/SOEs; is responsible for the review of nonconforming material and aids quality in determining its disposition; is responsible for change management process and maintaining the BaaN Item Masters.

- The Quality Department is responsible for verification of incoming material, analysis of data in the test and inspection areas for manufacturing continuous process improvement, lean manufacturing initiatives and administering the internal audit program and corrective/preventive action programs.  They and Purchasing/Sourcing are jointly responsible for the quality of purchased material.

- Purchasing/Sourcing is jointly responsible for the quality of purchased material; responsible for the availability of purchased material; responsible for maintaining the approved supplier list.

- Operations and Logistics are responsible for all planning, material procurement, manufacturing of products, final testing and inspection of finished products, collection of data in the test and inspection areas for continuous process improvement, maintenance of buildings, grounds and equipment, packaging, storage, handling and shipping of product.

- The Customer Care Team / Customer Service is responsible for customer relations for post order service and are responsible for all supply parts replacement for warranty and non-warranty requirements, overseas and domestic.

- The Human Resources Department is responsible for making sure each job function has a job description and retaining personnel records. They are also involved in establishing employment rules for the AxisClinical organization.

- Safety Environmental Department  is responsible for initiating safety training, managing the environmental, health and safety activities as well as establishing safety procedures for the AxisClinical organization.

- The Information Systems Department is responsible for maintaining records of validation of new and existing software, maintaining the functionality of the BaaN system and backing up data.

- The Controller is responsible for making sure that Finance and Accounting processes are compliant with the Sarbanes Oxley Law; coordinating and working with other departments as necessary and working with Corporate to schedule appropriate audits.

- All employees at AxisClinical have the authority to identify nonconformities, control further processing or stop production processes until the problem has been corrected.

- Roles for regulatory requirements:

  - US FDA Official Correspondent – AxisClinical LLC – Quality Compliance Mgr. – Primary FDA Contact

  - CE Authorized representative – AxisClinical Electron GmbH, 1 Robert-Bosch Str, Langenselbold, 63505, Germany

  - China CFDA registrations – Regulatory Affairs Manager, Life Sciences China

  - Latin America Registrations – Sr Regulatory Affairs Manager, LATAM

  - South Korean Registration – Regulatory Affairs Manager, AxisClinical Korea

Responsibility and Interaction for QMS Processes